## Slide 1: Clinical trials for medical devices: FDA and the IDE process
Owen Faris, Ph.D.
owen.faris@fda.hhs.gov
Director, Clinical Trials Program
Office of Device Evaluation
Center for Devices and Radiological Health, FDA

## Slide 2: What is a Medical Device?
The Section 201(h) of the Food, Drugs and Cosmetics Act defines a medical device as any healthcare product that does not achieve its principal intended purposes by chemical action or by being metabolized.
  As simple as a tongue depressor or a thermometer
  As complex robotic surgery devices
©2006 Intuitive Surgical, Inc.

## Slide 3: Device Classification
Medical Device Classes
Class I
  General Controls
  Most exempt from premarket submission
Class II
  Special Controls
  Premarket Notification [510(k)]
Class III
  Premarket Approval
  Require Premarket Application [PMA]

## Slide 4: 510(k) Premarket Notification
Substantial equivalence
10-15% require clinical data
Performance testing
Usually confirmatory
Type of study dictated by:
  Ability of bench and animal testing to answer questions
  Amount of difference between subject device and predicate

## Slide 5: PMA Premarket Approval Application
Establish reasonable assurance of safety and effectiveness
Bench-Animal-Human
Clinical Studies
  Feasibility and pivotal

## Slide 6: Stages of review for PMA device
Pre-Sub
IDE
PMA
PMA-S
Discuss:
Device design
Bench testing
Animal testing
Clinical trial
Request
approval for
clinical trial
Request
market
approval
Request
approval for
device change
or upgrade
(which may
require a new
IDE)

## Slide 7: Today’s focus:
Pre-Sub
IDE
PMA
PMA-S
Discuss:
Device design
Bench testing
Animal testing
Clinical trial
Request
approval for
clinical trial
Request
market
approval
Request
approval for
device change
or upgrade
(which may
require a new
IDE)

## Slide 8: What is an Investigational Device Exemption (IDE)?
FDA approval of an IDE is required for US human study of a significant risk device which is not approved for the indication being studied.

## Slide 9: Device trials are unique
Trials tend to be smaller than drug trials
Some novel, many “me-too”
Many difficult to blind, randomize, control
Many depend on physician technique
Device modifications occur during trial
Endpoints highly diverse
Typically, single pivotal trial follows feasibility stage(s)
Designed to support a “reasonable assurance of safety and effectiveness” for the marketing application

## Slide 10: Types of IDEs
Feasibility study
  May provide support for a future pivotal study or may be used to answer basic research questions
  Not intended to be the primary support for a marketing application
  Endpoints and sample size generally not statistically driven
  Often required by FDA prior to pivotal study to assess basic safety and potential for effectiveness
  Generally ~10-40 patients but may be larger
  FDA review is primarily focused on safety and whether the potential benefit or value of the data justifies risk

## Slide 11: Types of IDEs
Early Feasibility Study (EFS)
  Generally <15 subjects
  Device may be early in development, before the device design has been finalized
  Approval of an early feasibility study IDE may be based on less nonclinical data
  Some nonclinical testing may be deferred
  Additional mitigations may be needed to protect subjects
  New EFS program within CDRH is designed to facilitate and support EFS research under IDE

## Slide 12: Types of IDEs
Pivotal study
  Generally intended as the primary clinical support for a marketing application
  Designed to demonstrate a “reasonable assurance of safety and effectiveness”
  Endpoints and sample size statistically driven
  Designed to assess both safety and effectiveness
  FDA review is much more complex

## Slide 13: FDA’s Feasibility IDE Review
Focused on safety
Critical issues
  Reasonable study conceptually?
  Adequate preclinical validation of device?
    Why is clinical really the next necessary step?
  Appropriate mitigation of potential risks?
  Appropriate enrollment criteria?
  Patients adequately informed?
  Sample size appropriate?

## Slide 14: FDA’s Pivotal IDE Review
Focused on safety and plan for collecting and evaluating study data
Additional critical issues
  Trial endpoints
  Randomization, blinding, follow-up, etc
  Study conduct and monitoring
  Statistical analysis plan

## Slide 15: Basic Submission Elements
Background of medical issue, the study goals, and why this study will further the science
Detailed description of the device under study
Previous studies (preclinical and clinical)
  Summary of available data
  Why is a clinical study needed at this stage?
  What evidence supports the safety of this study/device and the potential for the study data to be meaningful?
  Are there outstanding safety questions that should be addressed with preclinical data?

## Slide 16: Basic Submission Elements
Risk analysis
  What are the potential risks to the patient?
  Does the study mitigate the risks where possible?
  Are the risks outweighed by the potential for benefit and/or value of the study
Patient monitoring and follow-up plan
Inclusion and exclusion criteria
Informed consent document
Sample size and number of investigational centers, with justification

## Slide 17: Submission Elements, Pivotal IDEs
Primary and secondary endpoints
  Discussion of appropriateness of endpoint parameters, hypotheses, and success criteria
Basic trial design
  Controlled? If not, why not?
  Randomized? If not, why not?
  Blinded? If not, why not?

## Slide 18: Submission Elements, Pivotal IDEs
Trial conduct and study monitoring
  Data handling and adjudication process
  Sponsor blinding
  Independent committees
  Case report forms
    Is the right information being gathered to support the study endpoints and are investigators adequately prompted to report adverse events?

## Slide 19: Submission Elements, Pivotal IDEs
Statistical analysis plan
  Clearly defined S & E hypotheses
  Type-1 error and multiplicity
  Missing data handling
  Sample size calculations and assumptions
  Assessment of critical covariates
  Adaptive design plans
  Interim analyses and early stopping rules
  Data handling

## Slide 20: Primary Endpoint Design
Should evaluate the safety and effectiveness of the device in the population expected to be indicated.
Generally divided into
  1 or more “safety” endpoints
  1 or more “effectiveness” endpoints
A study would be considered successful if both the safety and effectiveness endpoints are met.

## Slide 21: Primary Endpoint Design
The clinical protocol should clearly and prospectively detail:
  Methods for obtaining endpoint data
  Definitions for what will be counted as a primary event in the analysis
  Situations in which patient data will be excluded
  How missing data will be handled
  How the impact of covariates will be assessed

## Slide 22: Sample Size & Follow-Up
Driven by either:
  Primary safety endpoint
  Primary effectiveness endpoint
Minimum number of patients and/or minimum duration of follow-up may be required depending on:
  Understanding of the safety and effectiveness of the device
  Concerns regarding durability of device safety or effectiveness

## Slide 23: Secondary Endpoints
Generally used to evaluate additional meaningful claims
Generally only considered if primary endpoints are successful
Should be used to provide further insight into the device effects and mechanisms of action
Definitions and analysis methods should be clearly detailed prospectively
Not considered “statistically significant” unless a pre-specified alpha allocation plan is in the protocol, even if the p-value is < 0.05

## Slide 24: Submission Elements, Pivotal IDEs
  Provide enough detail to avoid ambiguity once the trial has started.

## Slide 25: Best Practices
Organization
Device Description
Basic Safety and Essential Performance
Investigational Plan
Interactions with FDA
Cohesive Narrative
Notes: Q-Sub guidance is 57 pages
EFS guidance is 40 pages
If you search for “clinical” in the guidance database, narrowing your focus to CDRH, you will find 58 guidance documents. Only 1 of those documents, the EFS guidance, is focused on the content of IDE submissions.

## Slide 26: FDA’s IDE Review Decisions
Approval
  Approves the trial for a specified number of patients and investigational centers
Approval with Conditions
  Allows sponsor to begin the trial if the sponsor agrees to address the conditions (deficiencies) from the conditional approval letter within 45 days
Disapproval
  Trial may not start until sponsor addresses the deficiencies from the letter, submits this information to FDA, and receives approval

## Slide 27: Significant Improvements
Median Days to Full Approval
442 days (2011) to 30 days (2016)
Percent Fully Approved within Two Cycles
~15% (2011) to ~71% (2016)
61% of IDE’s Approved in 1st Round!
Notes: CDRH review teams (and sponsors) have done an amazing job, significantly reducing the number of cycles and time to full approval
Primary factor improving these results is interacting with the sponsor during and after each review cycle
IDE submission quality varies widely
Interactions are stressful and resource intensive
Some IDE submissions aren’t organized well and lack critical information: we expend tremendous resources on these submissions
Incorporating a refuse-to-accept process would make the review process more stressful and require even more resources
Stressful for review staff AND sponsors

## Slide 28: Does study failure imply PMA disapproval?
Sometimes but not always.
PMA approval is based on a Benefit-Risk assessment
FDA is always willing to review all available data to determine whether there is a reasonable assurance that the device safe and effective.

## Slide 29: Does study failure imply device disapproval?
Alternatives
  Unexpected safety concerns are outweighed by stronger than expected benefit
  Inconclusive study result is supplemented by other clinical or non-clinical data
  Device is safe and effective for some limited indication or patient population
  All of these alternatives may raise serious type-1 error concerns. FDA is therefore very conservative in its consideration of these alternatives.

## Slide 30: Does study success imply device approval?
Often but not always
Sometimes the primary endpoints do not capture a serious unexpected safety concern that is observed in the trial.
Other clinical or non-clinical data may conflict with the study result.
Can result in:
  Device disapproval
  Requirement for more data
  Limited indication

## Slide 31: Conclusions
One size does not fit all for device trials
Pivotal studies should be designed to evaluate whether there is a “reasonable assurance of safety and effectiveness.”
PMA approvability is based upon a Benefit-Risk assessment which strongly considers outcome of primary safety and effectiveness endpoints.

## Slide 32: Conclusions
Secondary endpoints are generally used to support claims if the primary endpoints are successful.
All endpoint analyses and definitions should be clearly pre-specified in the approved clinical protocol.
Trial design is challenging. We recommend talking to FDA early through the pre-submission process.

## Slide 33: Our Shared Goal
For Sponsors
To make the process of preparing the IDE and interacting with FDA to fulfill the IDE regulations and to obtain full approval as clear, quick, and predictable as possible
as possible
For FDA Staff
To make the process of reviewing the IDE and interacting with sponsors as efficient, focused, and consistent as possible

## Slide 34: Important Guidance Documents
FDA Decisions for IDE Clinical Investigations
  http://www.fda.gov/downloads/MedicalDevices/DeviceRegulationandGuidance/GuidanceDocuments/UCM279107.pdf
Design Considerations for Pivotal Clinical Investigations for Medical Devices
  http://www.fda.gov/downloads/MedicalDevices/DeviceRegulationandGuidance/GuidanceDocuments/UCM373766.pdf
IDEs for Early Feasibility Medical Device Clinical Studies
  http://www.fda.gov/downloads/medicaldevices/deviceregulationandguidance/guidancedocuments/ucm279103.pdf
Evaluation of sex-specific data in medical device clinical studies
  http://www.fda.gov/downloads/MedicalDevices/DeviceRegulationandGuidance/GuidanceDocuments/UCM283707.pdf

## Slide 35: Online Resources
CDRH Learn – Online Regulatory Training Tool
  Over 50 Medical device and Radiological Health modules
  Video and PowerPoint presentations available 24/7
  Certificate of completion upon passing post-tests
  Many modules are translated into Chinese and Spanish
  http://www.fda.gov/Training/CDRHLearn/
Device Advice – Online Regulatory Information
  Searchable by topic
  http://www.fda.gov/MedicalDevices/DeviceRegulationandGuidance/
Division of Industry and Consumer Education (DICE)
  1-800-638-2041
  301-796-7100
  dice@fda.hhs.gov